## Slide 1: Computed Tomography
CSE 5780 Medical Imaging Systems and Signals
Ehsan Ali and Guy Hoenig

## Slide 2: Computed Tomography using ionising radiations
Medical imaging has come a long way since 1895 when Röntgen first described a ‘new kind of ray’.
That X-rays could be used to display anatomical features on a photographic plate was of immediate interest to the medical community at the time.
Today a scan can refer to any one of a number of medical-imaging techniques used for diagnosis and treatment.

## Slide 3: Instrumentation (Digital Systems)
The transmission and detection of X-rays still lies at the heart of radiography, angiography, fluoroscopy and conventional mammography examinations.
However, traditional film-based scanners are gradually being replaced by digital systems
The end result is the data can be viewed, moved and stored without a single piece of film ever being exposed.

## Slide 4: CT Imaging
Goal of x-ray CT is to reconstruct an image whose signal intensity at every point in region imaged is proportional to μ (x, y, z), where μ is linear attenuation coefficient for x-rays.
In practice, μ is a function of x-ray energy as well as position and this introduces a number of complications that we will not investigate here.
X-ray CT is now a mature (though still rapidly developing) technology and a vital component of hospital diagnosis.

## Slide 5: Comparisons of CT Generations
Comparison of CT Generations |  |  |  |  |  |  | 
Generation | Source | Source Collimation | Detector | Detector Collimation | Source-Detector Movement | Advantages | Disadvantages
1G | Single x-ray tube | Pencil beam | Single | None | Move linearly and rotate in unison | Scattered energy is undetected | Slow
2G | Single x-ray tube | Fan beam, not enough to cover FOV | Multiple | Collimated to source direction | Move linearly and rotate in unison | Faster than 1G | Lower efficiency and larger noise because of the collimators in directors
3G | Single x-ray tube | Fan beam, enough to cover FOV | Many | Collimated to source direction | Rotate in synchrony | Faster than 2G, continuous rotation using slip ring | Moe expensive than 2G, low efficiency
4G | Single x-ray tube | Fan beam covers FOV | Stationary ring of detectors | Cannot collimate detectors | Detectors are fixed, source rotates | Higher efficiency than 3G | High scattering since detectors are not collimated
5G (EBCT) | Many Tungsten anodes in a single large tube | Fan beam | Stationary ring of detectors | Cannot collimate detectors | No moving parts | Extremely fast, capable of stop-action imaging of beating heart | High cost, difficult to calibrate
6G (Spiral CT) | 3G/4G | 3G/4G | 3G/4G | 3G/4G | 3G/4G plus linear patient table motion | Fast 3D images | A bit more expensive
7G (Multi-slice CT) | Single x-ray tube | Cone beam | Multiple arrays of detectors | Collimated to source direction | 3G/4G/6G motion | Fast 3D images | Expensive

## Slide 6: Four generations of CT scanner

## Slide 7: X-rays CT - 1st Generation
Single X-ray Pencil Beam
Single (1-D) Detector set at 180 degrees opposed
Simplest & cheapest scanner type but very slow due to
Translate(160 steps)
Rotate (1 degree)
~ 5minutes (EMI CT1000)
Higher dose than fan-beam scanners
Scanners required head to be surrounded by water bag

## Slide 8: Fig 1: Schematic diagram of a 1st generation CT scanner
(a) X-ray source projects a thin “pencil” beam of x-rays through sample, detected on the other side of the sample. Source and detector move in tandem along a gantry. (b) Whole gantry rotates, allowing projection data to be acquired at different angles.

## Slide 9: First Generation CT Scanner

## Slide 10: First-generation CT Apparatus
Notes: Intensity input @ position ‘x’= Intensity output (exp {MU(x,y)dy}

## Slide 11: Further generations of CT scanner
The first-generation scanner described earlier is capable of producing high-quality images. However, since the x-ray beam must be translated across the sample for each projection, the method is intrinsically slow.
Many refinements have been made over the years, the main function of which is to dramatically increase the speed of data acquisition.

## Slide 12: Further generations of CT scanner (cont’d)
Scanner using different types of radiation (e.g., fan beam) and different detection (e.g., many parallel strips of detectors) are known as different generations of X-ray CT scanner. We will not go into details here but provide only an overview of the key developments.

## Slide 13: Second Generation CT scanner

## Slide 14: X-rays CT - 2nd Generation (~1980)
Narrow Fan Beam X-Ray
Small area (2-D) detector
Fan beam does not cover full body, so limited translation still required
Fan beam increases rotation step to ~10 degrees
Faster (~20 secs/slice) and lower dose
Stability ensured by each detector seeing non-attenuated x-ray beam during scan

## Slide 15: Third Generation CT Scanner

## Slide 16: X-rays CT - 3rd Generation (~1985)

## Slide 17: X-rays CT - 3rd Generation (~1985)

## Slide 18: X-rays CT - 3rd Generation
Wide-Angle Fan-Beam X-Ray
Large area (2-D) detector
Rotation Only - beam covers entire scan area
Even faster (~5 sec/slice) and even lower dose
Need very stable detectors, as some detectors are always attenuated
  Xenon gas detectors offer best stability (and are inherently focussed, reducing scatter)
  Solid State Detectors are more sensitive - can lead to dose savings of up to 40% - but at the risk of ring artefacts

## Slide 19: X-rays CT - 3rd Generation Spiral

## Slide 20: X-rays CT – 3rd Generation Multi Slice
Latest Developments - Spiral, multislice CT Cardiac CT

## Slide 21: X-rays CT – 3rd Generation Multi Slice

## Slide 22: Fourth Generation CT Scanners

## Slide 23: X-rays CT - 4th Generation (~1990)
Wide-Angle Fan-Beam X-Ray: Rotation Only
Complete 360 degree detector ring (Solid State - non-focussed, so scatter is removed post-acquisition mathematically)
Even faster (~1 sec/slice) and even lower dose
Not widely used – difficult to stabilise rotation + expensive detector
Fastest scanner employs electron beam, fired at ring of anode targets around patient to generate x-rays.
Slice acquired in ~10ms - excellent for cardiac work
X-rays CT - Electron Beam 4th Generation

## Slide 24: X-Ray Source and Collimation

## Slide 25: CT Data Acquisition

## Slide 26: CT Detectors: Detector Type

## Slide 27: Xenon Detectors

## Slide 28: Ceramic Scintillators

## Slide 29: CT Scanner Construction: Gantry, Slip Ring, and Patient Table

## Slide 30: Reconstruction of CT Images: Image Formation
REFERENCE DETECTOR
REFERENCE DETECTOR
ADC
PREPROCESSOR
COMPUTER
RAW DATA
CONVOLVED DATA
BACK
PROJECTOR
RECONSTRUCTED DATA
PROCESSORS
DISK
TAPE
DAC
CRT DISPLAY

## Slide 31: The Radon transformation
In a first-generation scanner, the source-detector track can rotate around the sample, as shown in Fig 1. We will denote the “x-axis” along which the assembly slides when the assembly is at angle φ by xφ and the perpendicular axis by yφ.
Clearly, we may relate our (xφ, yφ) coordinates to the coordinates in the un-rotated lab frame by [5]

## Slide 32: Figure 2: Relationship between Real Space and Radon Space
Highlighted point on right shows where the value λφ (xφ) created by passing the x-ray beam through the sample at angle φ and point xφ is placed. Note that, as is conventional, the range of φ is [-π / 2, +π / 2], since the remaining values of φ simply duplicate this range in the ideal case.

## Slide 33
Hence, the “projection signal” when the gantry is at angle φ is
[6]
We define the Radon transform as
[7]

## Slide 34: Attenuation (x-ray intensity)

## Slide 35
X-ray Attenuation (cont’d)

## Slide 36: Radon Space
We define a new “space”, called Radon space, in much the same way as one defines reciprocal domains in a 2-D Fourier transform. Radon space has two dimensions xφ and φ . At the general point (xφ, φ), we “store” the result of the projection λφ(xφ).
Taking lots of projections at a complete range of xφ and φ “fills” Radon space with data, in much the same way that we filled Fourier space with our 2-D MRI data.

## Slide 37: CT ‘X’ Axis
‘X’
Axis

## Slide 38: CT ‘Y’ Axis
‘Y’
Axis

## Slide 39: CT ‘Z’ Axis
‘Z’
Axis

## Slide 40: CT Isocenter
ISOCENTER

## Slide 41: Fig 3. Sinograms for sample consisting of a small number of isolated objects.
In this diagram, the full range of φ is [-π, +π ] is displayed.

## Slide 42: Relationship between “real space” and Radon space
Consider how the sinogram for a sample consisting of a single point in real (image) space will manifest in Radon space.
For a given angle φ, all locations xφ lead to λφ(xφ) = 0, except the one coinciding with the projection that goes through point (x0, y0) in real space. From Equation 5, this will be the projection where xφ = x0 cos φ + y0 sin φ.

## Slide 43
Thus, all points in the Radon space corresponding to the single-point object are zero, except along the track
[8]
where R = (x2 + y2)1/2 and φ0 = tan-1 ( y / x).
If we have a composite object, then the filled Radon space is simply the sum of all the individual points making up the object (i.e. multiple sinusoids, with different values of R and φ0). See Fig 3 for an illustration of this.

## Slide 44: Reconstruction of CT images (cont’d)
This is performed by a process known as back-projection, for which the procedure is as follows:
Consider one row of the sinogram, corresponding to angle φ. Note how in Fig 3, the value of the Radon transform λφ(xφ) is represented by the grey level of the pixel. When we look at a single row (i.e., a 1-D set of data), we can draw this as a graph — see Fig 4(a). Fig 4(b) shows a typical set of such line profiles at different projection angles.

## Slide 45: Fig 4a. Relationship of 1-D projection through the sample and row in sinogram

## Slide 46: Fig 4b. Projections at different angles correspond to different rows of the sinogram

## Slide 47: Fig 4c. Back-projection of sinogram rows to form an image. The high-intensity areas of image correspond to crossing points of all three back-projections of profiles.

## Slide 48
General Principles
of Image Reconstruction
Image Display - Pixels and voxels

## Slide 49: PIXEL Size Dependencies:
MATRIX SIZE
FOV

## Slide 50: PIXEL vs VOXEL
PIXEL
VOXEL

## Slide 51: VOXEL Size Dependencies
FOV
MATRIX SIZE
SLICE THICKNESS

## Slide 52: Pixel MATRIX

## Slide 53: Reconstruction Concept
 |  | Ц |  | 
 |  | CT# |  | 
RECONSTRUCTION

## Slide 54: CT and corresponding pixels in image

## Slide 55
Simple numerical
example

## Slide 56: µ To CT number

## Slide 57: CT Number Flexibility
We can change the appearance of the image by varying the window width (WW) and window level (WL)
This spreads a small range of CT numbers over a large range of grayscale values
This makes it easy to detect very small changes in CT number

## Slide 58: Windowing in CT

## Slide 59: CT Numbers
Linear attenuation coefficient of each tissue pixel is compared with that of water:

## Slide 60: CT Number Window

## Slide 62
Example values of μt:
At 80 keV: μbone = 0.38 cm-1
μwater = 0.19 cm-1
The multiplier 1000 ensures that the CT (or Hounsfield) numbers are whole numbers.

## Slide 63: Linear Attenuation Coefficient ( cm-1)
BONE 0.528
BLOOD 0.208
G. MATTER 0.212
W. MATTER 0.213
CSF 0.207
WATER 0.206
FAT 0.185
AIR 0.0004

## Slide 64: CT # versus Brightness Level
+ 1000
-1000

## Slide 65: CT in practice

## Slide 66: SCAN Field Of View (FOV) Resolution
SFOV

## Slide 67: Display FOV versus Scanning FOV
DFOV CAN BE EQUAL OR LESS OF SFOV
SFOV – AREA OF MEASUREMENT DURING SCAN
DFOV - DISPLAYED IMAGE

## Slide 68: Image Quality in CT

## Slide 69: Projections

## Slide 70: Back Projection
Reverse the process of measurement of projection data to reconstruct an image
Each projection is ‘smeared’ back across the reconstructed image

## Slide 71: Back Projection

## Slide 72: Filtered Back Projection
Back projection produces blurred trans-axial images
Projection data needs to be filtered before reconstruction
Different filters can be applied for different diagnostic purposes
  Smoother filters for viewing soft tissue
  Sharp filters for high resolution imaging
Back projection process same as before

## Slide 73: Filtered Back Projection

## Slide 74: Filtered Back Projection

## Slide 75: Filtered Back Projection

## Slide 76: Summary and Key Points
A tomogram is an image of a cross-sectional plane or slice within or through the body
X-Ray computed tomography (CT) produces tomograms of the distribution of linear attenuation coefficients, expressed in Hounsfield units.
There are currently 7 generations of CT scanner design, which depend on the relation between the x-ray source and detectors, and the extent and motion of the detectors (and patient bed).
The basic imaging equation is identical to that for projection radiography; the difference is that the ensemble or projections is used to reconstruct cross-sectional images.
The most common reconstruction algorithm is filtered back projection, which arises from the projection slice theorem.
In practice, the reconstruction algorithm must consider the geometry of the scanner–parallel-beam, fan-beam,helical-scan, or cone-beam.
As in projection radiography, noise limits an image’s signal to noise ratio.
Other artifacts include aliasing , beam hardening, and – as in projection radiography – inclusion of the Compton scattered photons.

## Slide 77: Cone Beam CT: Introduction

## Slide 78: CT Basic Principle
Point 1: Purpose of CT and Basic principle
Point 2: The internal structure of an object can be reconstructed from multiple projections of the object
Point 3: Computerized Tomography, or CT is the preferred current technology

## Slide 79: The Basics

## Slide 80: Current Cone Beam Systems

## Slide 81: Cone Beam Reconstruction

## Slide 82: The ‘Z’ Axis

## Slide 83: Medical CT Vs. Cone Beam CT

## Slide 84: Medical CT Example

## Slide 85: Cone Beam CT Example

## Slide 86: Cone Beam CT example

## Slide 87: CBCT Advantages over Medical CT

## Slide 88: References
University of Surrey: PH3-MI (Medical Imaging): David Bradley Office 18BC04 Extension 3771
Physical Principles of Computed Tomography
Basic Principles of CT Scanning: ImPACT Course October 2005