## Slide 1: Physics of Ultrasound
Krystal Kerney
Kyle Fontaine
Ryan O’Flaherty

## Slide 2: Basics of Ultrasound
Ultrasound is sound with frequencies higher than about 20 kHz
  For medical ultrasound, systems operate at much higher frequencies, typically 1 – 10 MHz
Propagation of ultrasound waves are defined by the theory of acoustics
  Ultrasound moves in a wavelike fashion by expansion and compression of the medium through which it travels
  Ultrasound waves travel at different speeds depending on material
  Ultrasound waves can be absorbed, refracted, focused, reflected, and scattered.

## Slide 3: Basics of Ultrasound
Process Overview
  Transducer (electrical signal a acoustic signal) generates pulses of ultrasound and sends them into patient
  Organ boundaries and complex tissues produces echoes (reflection or scattering) which are detected by the transducer
  Echoes displayed on a grayscale anatomical image
    Each point in the image corresponds to an anatomical location of an echo-generating structure
    Brightness corresponds to echo strength

## Slide 4: Wave Equation
Acoustic wave
  Pressure wave that propagates through material via compression and expansion
    Compress a small volume of tissue
    Releasing it causes it to expand past equilibrium
    Surrounding tissues are compressed, sequence starts again
In soft tissue, particles oscillate in same direction as wave
  This is a longitudinal wave
Speed of sound (c) waves is dependent on compressibility (k) and density (ρ)
Table 10.1
  Tissue ≈ 1540 m/s
  Air ≈ 330 m/s

## Slide 5: Wave Equation
Acoustic wave can be described as spatially dependent, time-varying pressure function
  Acoustic Pressure P ( x , y , z , t )
For longitudinal waves: P = Zv
  Z = cρ, characteristic impedance
  v, particle speed, generally NOT equivalent to speed of sound (c)
    Acoustic pressure (P = Zv) analogous to electrical circuits (V = IR)
    Acoustic pressure waves must satisfy the following 3-D wave equation
  Where is the 3-D Laplacian operator

## Slide 6: Wave Equation – Plane Waves
Aforementioned equation is hard to solve, simplify by considering 2 special cases
  Plane Waves
  Spherical Waves
Plane waves vary only in one spatial direction and time
  Consider a plane wave moving in the +z or –z direction
  P ( z , t ) = P ( x , y , z , t )
  Plugging this into the 3-D wave equation yields the 1-D wave equation
  General Solution :
    is a forward traveling wave, is a backward traveling wave
  Later we will approximate acoustic waves from certain transducers as plane waves

## Slide 7: Wave Equation – Plane Waves
An important aside…
  The sinusoidal function satisfies the 1-D wave equation
  =
    Hold z fixed, pressure around a fixed particle varies sinusoidally with radial frequency of = kc
      f = = with units of cycles per second or Hz
    Hold t fixed, the pressure at a particular time varies sinusoidally with radial spatial frequency k, the wave number
    Wavelength λ = with units of length
    Alternate form yields important relationship between wavelength, speed of sound, and frequency  λ =

## Slide 8: Wave Equation – Spherical Waves
Spherical waves depend on only time and the radius from the source of disturbance.
Can be generated in an isotropic material via a small local disturbance in pressure
  with the source at (0,0,0)
  Realizing that P = ( r , t ) and noting r as a function of x, y, and z, we can rearrange the 3-D wave equation
    , the spherical wave equation
  General solution :
    Where is an outward traveling wave, and is an inward traveling wave (generally inward traveling waves don’t exist)
  Hence :
    Similar to forward traveling wave with additional loss factor of 1/r as it travels radially out and loses amplitude due to increasing surface area

## Slide 9: Wave Propagation – Acoustic Energy and Intensity
Acoustic waves carry energy with them
  Particles in motion have kinetic energy
    wk = ½ρv2
  Particles prepared to move have potential energy
    wp = ½κp2
  Acoustic energy density is defined by the sum of the kinetic energy density and the potential energy density
    w = wk + wp
  Acoustic Intensity
    I = pv
    Also called the acoustic energy flux
  Acoustic energy density and acoustic intensity are related via the equation of energy conservation

## Slide 10: Wave Propagation – Reflection and Refraction at Plane Interfaces
See Figure 10.2
This is called Snell’s Law
If solving for and ; does not exist
  Conclude that all energy is reflected
If all incident angles above critical angle will result in total reflection
  for

## Slide 11: Wave Propagation – Transmission and Reflection Coefficients
Since incident, reflected, and transmitted waves all meet at the interface, the tangential particle motion caused by the incident wave must coincide with the sum of the tangential particle motions of transmitted and reflected waves
  If you plug in for acoustic pressure and acoustic intensity and consider that pressure is continuous across the interface you can find:
    Pressure reflectivity, and intensity reflectivity
      ,
    Pressure transmittivity, and intensity transmittivity
      ,

## Slide 12: Wave Propagation – Attenuation
Attenuation – accounts for loss of wave amplitude due to all mechanisms, including absorption, scattering, and mode conversion.
  Absorption is the process by which wave energy is converted to thermal energy then dissipated into the medium.
  Scattering is the process by which secondary spherical waves are generated as the wave propagates.
  Mode conversion is the process by which longitudinal waves are converted to transverse shear waves (and vice versa).

## Slide 13: Wave Propagation - Attenuation
A forward-traveling plane wave with attenuation:
Amplitude decay
Amplitude attenuation factor
Phenomelological – agrees well in practice but is not easily supported by theory
If given an attenuation coefficent, , convert to then you can use the forward-traveling plane wave with attenuation equation.

## Slide 14: Wave Propagation - Attenuation
When attenuation is only due to the conversion of acoustic energy to thermal energy, the attenuation coefficient is called the absorption coefficent.

## Slide 15: Wave Propagation - Scattering
Many targets within the body are significantly smaller than the acoustic wavelength.
Under these circumstances, assume that when the target is excited by an incident acoustic plane wave, it vibrates as a small spherical body, which gives rise to spherical waves.
See Figure 10.3
Small target at (0,0,d) acts as a spherical wave source, converts a fraction of the incident wave into a spherical wave
  The fraction of the incident wave converted into the spherical wave is denoted R, and is the reflection coefficient. It is a property of the individual target and the embedding medium.
  Scattered wave equation

## Slide 16: The Doppler Effect
Change in frequency of the sound due to the relative motion of the source and/or receiver
  Example: ambulance
  Figure 10.4 (a)
    , but the frequency shift is only dependent on the component of source velocity in the direction of the observer
    where ϴ is angle between the vector pointing from source to receiver and the vector pointing from source to direction of motion

## Slide 17: The Doppler Effect
The difference between observed and source frequencies is the Doppler frequency
  Thus, , but because c >> v,
  Sign of Doppler frequency indicates direction
    + : Source moving towards observer
    - : Source moving away from observer
  Pulse echo mode: Transducer is both source of sound and receiver of the Doppler-shifted echo returning from the object.
    Sound collected by transducer is received by moving object and retransmitted by moving object.

## Slide 18: The Doppler Effect
Consider 10.4 (b)
  T is stationary source, O is moving receiver.
  Moving object observes a frequency, [*]
  And the corresponding Doppler frequency,
  Equivalent to
In pulse echo mode, the echo received by T will be shifted by both the effects of a moving receiver and a moving source
  Essentially 2x the Doppler Frequency than in either case alone

## Slide 19: The Doppler Effect
Transducer T generates wave with frequency fS
Object O, moving with velocity v at angle relative to u, recieves a frequency fO.
  This frequency was shown before as [*]
The object reflects or scatters the wave, so it is now a moving source with frequency fO
The stationary transducer now receives a frequency fT
  =
The Doppler frequency in pulse echo mode will reduce to
Pg. 332, final paragraph
  Doppler shift velocimeter
  |fD| evaluation - Doppler Motion
  fD magnitude and sign - Doppler Imagaing

## Slide 20: Beam Pattern Formation
Simple Field Pattern Model
  Geometric Approximation
  Fresnel Region
  Fraunhofer Region
Wave diameter

## Slide 21: Beam Pattern Formation
Diffraction Formulation
  Narrow Bandpulse Model
  Where ‘n’ is the pressure signal and ‘’ is the envelope (Fig 10.6)

## Slide 22: Beam Pattern Formation
Received Signal with Field Pattern
  See Fig 10.7, assume a spherical scatter at (x,y,z)
  Pressure Summation:
  Received Electrical Waveform
  These equations are to be simplified using several assumptions
    Plane wave Approximation
    Paraxial Approximation
    Fresnel Approximation
    Fraunhofer Approximation
  See Eq. 10.65,.66,.72

## Slide 23: Focusing
Works through:
  Electrical means
  Geometric Adjustment to the transducer crystal
  Applying a lens
Curved Lenses or Vibrators focus sound in the same way that convex optical lenses focus light.
Increased resolution at the focal depth comes at the cost of range.

## Slide 24: Ultrasound Imaging Systems
Krystal Kerney
Kyle Fontaine
Ryan O’Flaherty

## Slide 25: Introduction
First, do no harm
Poses no known risk to the patient
Least expensive tool for the job
Portable (necessary to move from bedside to operating room)

## Slide 26: Instrumentation
Ultrasound Transducer
  Transducer Materials
  Resonance
Ultrasound Probes
  Single- Element Probes
  Mechanical Scanners
  Electronic Scanners

## Slide 27: Ultrasound Transducers
Transducer Materials
  Piezoelectric Crystals – translates mechanical strain into electrical signal and vice versa
  Most common material is Lead Zirconate Titanate (PZT)
  Selected for a high d and g constants
  Transmitting constant, d, relates strain to a unit electric field
  Receiving constant, g, relates potential produced by unit stress
  Other materials include: Quartz, Polyvinylidene Fluoride (PVDF)

## Slide 28: Ultrasound Transducers
Resonance
  Crystals tend to vibrate sinusoidally after initial excitation due to incoming acoustic waves reflecting off the back end of the crystal
  Fundamental resonant frequency (FRF) represents when the reflected wave interferes constructively with its source.
    Where:
      is the speed of sound in the transducer
      is the thickness of the transducer
      is the Fundamental Resonant Frequency
      is the wavelength at the FRF

## Slide 29: Ultrasound Transducers
Medical Transducers tend to be ‘shock excited’
  This refers to their output behaving as an impulse
Once excited the in-transducer wave continues to resonate until it loses energy
  This energy is damped away using epoxy backing in the transducer with a high coefficient of absorption. This compensates for PZT’s low absorption coefficient and the high reflectivity between the body and transducer.
  This epoxy must have a similar impedance to PZT in order to maintain a low coefficient of reflectivity between them.
  The epoxy finishes damping away the in-transducer wave’s energy in approximately 3-5 cycles.

## Slide 30: Ultrasound Probes
Single Element Probes
  Simplest assembly of transducer
  Look to Figure 11.5, this illustrates the construction of a single element probe
  Lens or curved crystal
  Ultrasound beam requires steering
  Modern systems of scanning allow for real-time imaging

## Slide 31: Ultrasound Probes
Mechanical Scanners
  Rocking or rotating a transducer crystal or set of crystals
  Figure 11.6
  Rocker – transducer travels through the same sector in a repeating fashion, first clockwise then counterclockwise
  Rotating – transducer is switched in as it enters the sector – always counterclockwise
    Regardless of design, field of view is always shaped like a slice of pie

## Slide 32: Ultrasound Probes
Electronic Scanners
  Arrangement of elements in the assemblies is linear
  Each element is rectangular
  Focused using a lens
  Linear array probe
    Elements have widths on the order of a wavelength and are electronically grouped together making several elements appear as one.
  Phased array probe
    Elements have widths of a quarter wavelength and the timing of firing of the elements are electronically controlled in order to steer and focus the beam.

## Slide 33: Pulse-Echo Imaging
Two important augmentations to basic imaging
  Phased arrays
  Doppler imaging
Ideal ultrasound imaging system would reconstruct and display the spatial distribution of reflectivity
  Not possible!
    Transducer’s impulse response function blurs reflectivity
    Envelope detection creates artifacts called speckles

## Slide 34: Pulse-Echo Equation
General form of equation using Fresnel approximation and Fraunhofer approximation is shown in equation 11.2
TGC is time gain compensation to cancel the gain terms in equation 11.2
Users can manually adapt the system to more or less gain so that subtle features can be seen in images.
Inexpensive ultrasound systems use a simple envelope detection procedure as shown in figure 11.10
A-mode signal is equation 11.10, it is the fundamental signal in all of ultrasound imaging.

## Slide 35: Transducer Motion
To acquire images, the transducer must move
  Consider (x,y) plane
  Assume the transducer energy travels down a cylinder having the same shape as its’ face
  The envelope equation, when it becomes a function of time and space can be thought of as an estimate of the reflectivity function as a function of spatial position
  Look at equation 11.20 and subsequent paragraph

## Slide 36: Ultrasound Imaging Modes
A-Mode Scan
  Amplitude-mode signal
  Transducer is fired rapidly and a succession of signals can be displayed on an oscilloscope (See figure 11.11).
  The time between successive firings is called repetition time
    Interval should be long enough so that returning echoes have died out, but fast enough to capture motion
  Useful when looking at heart valve motion

## Slide 37: Ultrasound Imaging Modes
M-Mode Scan
  Using each A-mode signal as a column in an image
  Value of A-mode signal becomes the brightness of the M-mode image
  Motion is revealed by bright traces moving up and down across the image as shown in figure 11.12
  Most often used to image motion of heart valves and is therefore shown along with ECG.

## Slide 38: Ultrasound Imaging Modes
B-Mode Scan
  Created by scanning the transducer beam in a plane
    Example – moving the transducer in the x-direction while beam is aimed down the z-axis (figure 11.13)
  Succession of A-mode signals are keyed to the x-position of the transducer.
  Image is created by brightness-modulating a CRT along a column using the corresponding A-mode signal
  An advantage of manual-scan systems is that you can angle the transducer to hit the same point of the body from different directions
  When multiple views of the same tissue are included in a single B-mode image, it is referred to as compound B-mode scanning
    Disadvantage – suffer from severe artifacts due to refraction

## Slide 39: Ultrasound Imaging Modes
B-mode scanners
  Linear scanner – collection of transducers arranged in a line, does not require motion. Requires large flat area with which to maintain contact with the body.
    Abdominal imaging
    Obstetrics
  Mechanical sector scanner – pivots a transducer about an axis orthogonal to the transducer’s axis.
  Phased array sector-scanner – collection of very small transducer elements arranged in a line. Smaller than linear scanner. Advantage is that focus can be varied over time providing a dynamic focus. Disadvantage is that sidelobes of acoustic energy are generated and can lead to artifacts.

## Slide 40: Ultrasound Imaging Modes
Depth of penetration
  Limited by attenuation
  is total range wave can travel before attenuated below system threshold
  , depth of penetration can only be half of the above equation (round trip)
Pulse Repetition Rate
  New pulse generated only after echoes from previous pulse are gone
  , pulse repetition interval – round trip time to max depth of penetration
  , pulse repetition rate

## Slide 41: Ultrasound Imaging Modes
B-Mode Image Frame Rate
  If N pulses are required to generate an image  Frame rate
  Typical frame rates in commercial ultrasound systems around 10-100 frames/sec
  Low end values create a great deal of flicker, unacceptable
  Scan conversion solves this (converts polar to rectangular) by reading out data at higher rate.
  Can also reduce field of view to enable increased frame rate (reduce N)

## Slide 42: More on Imaging Modes
A-mode: A-mode (amplitude mode) is the simplest type of ultrasound. A single transducer scans a line through the body with the echoes plotted on screen as a function of depth. Therapeutic ultrasound aimed at a specific tumor or calculus is also A-mode, to allow for pinpoint accurate focus of the destructive wave energy.
B-mode or 2D mode: In B-mode (brightness mode) ultrasound, a linear array of transducers simultaneously scans a plane through the body that can be viewed as a two-dimensional image on screen. More commonly known as 2D mode now.
C-mode: A C-mode image is formed in a plane normal to a B-mode image. A gate that selects data from a specific depth from an A-mode line is used; then the transducer is moved in the 2D plane to sample the entire region at this fixed depth. When the transducer traverses the area in a spiral, an area of 100 cm2 can be scanned in around 10 seconds.
M-mode: In M-mode (motion mode) ultrasound, pulses are emitted in quick succession – each time, either an A-mode or B-mode image is taken. Over time, this is analogous to recording a video in ultrasound. As the organ boundaries that produce reflections move relative to the probe, this can be used to determine the velocity of specific organ structures.

## Slide 43: More on Imaging Modes
Doppler mode: This mode makes use of the Doppler effect in measuring and visualizing blood flow
  Color Doppler: Velocity information is presented as a color coded overlay on top of a B-mode image
  Continuous Doppler: Doppler information is sampled along a line through the body, and all velocities detected at each time point is presented (on a time line)
  Pulsed wave (PW) Doppler: Doppler information is sampled from only a small sample volume (defined in 2D image), and presented on a timeline
  Duplex: a common name for the simultaneous presentation of 2D and (usually) PW Doppler information. (Using modern ultrasound machines color Doppler is almost always also used, hence the alternative name Triplex.)
Pulse inversion mode: In this mode two successive pulses with opposite sign are emitted and then subtracted from each other. This implies that any linearly responding constituent will disappear while gases with non-linear compressibility stands out.
Harmonic mode: In this mode a deep penetrating fundamental frequency is emitted into the body and a harmonic overtone is detected. In this way depth penetration can be gained with improved lateral resolution.

## Slide 44: Steering and Focusing
Phased Array
Steering
  Add a separate delay element to each transducer to steer acoustic beam
  Firing times for each transducer to generate a plane wave in a known direction
  Eq. 11.31 and Figures 11.15, 11.16
Steering + Focusing
  Focusing = Refinement of steering
  Figure 11.17 and Eq. 11.34
    Delays are not multiples of a base delay
    Don’t need to be fire in same order as their geometric order

## Slide 45: Beamforming and Dynamic Focusing
Beamforming
  Plane wave incident upon the transducer from a direction ϴ will hit one transducer at the end of the array first and then successive transducers.
  Delay the received waveforms to coherently sum them  the entire ray is sensitized to direction ϴ
    Device delays are same as the transmit delays for steering + focusing
  Result is increased sensitivity for directions
  Figure 11.18
Dynamic Focusing
  Manipulates delays on transducer signals such that they have increased sensitivity to a particular point in space at a particular time.
  Figure 11.19, Eq. 11.39